## Phantom
Phantom | Weight (g)
2015-07-23 | 1988

## Depth
Depth | Axial Plane |  |  |  |  | Longitudinal Plane
| Freq (MHz) | Baseline (mm) | Current (mm) | Result | Image | Freq (MHz) | Baseline (mm) | Current (mm) | Result | Image
2015-08-11 | 6 | 80.18 | 74.16 | PASS | View Image | 6 | 78.8 | 66.14 | FAIL | View Image
Note on F3: Poon, Justin:
Note on K3: Poon, Justin:

## Axial Resolution
Axial Resolution | Axial Plane | [7 empty] | Longitudinal Plane
| Freq (MHz) |  | Baseline (mm) | Current (mm) | Diff (abs) | Diff (%) | Result | Image | Freq (MHz) |  | Baseline (mm) | Current (mm) | Diff (abs) | Diff (%) | Result | Image
2015-08-10 | 6 | Proximal (B1) | 0.46 | 0.58 | 0.12 | 23.7 | PASS | View Image | 6 | Filament 1 | 0.5 | 0.61 | 0.11 | 19.65 | PASS | View Image
|  | Proximal (F1) | 0.53 | 0.76 | 0.23 | 35.9 | PASS |  |  | Filament 6 | 0.48 | 0.49 | 0.01 | 2.86 | PASS
|  | Distal (B5) | 0.58 | 0.63 | 0.05 | 8.42 | PASS
|  | Distal (F5) | 0.62 | 0.65 | 0.03 | 4.14 | PASS
Note on I3: Poon, Justin:
Note on Q3: Poon, Justin:

## Lateral Resolution
Lateral Resolution | Axial Plane | [7 empty] | Longitudinal Plane
| Freq (MHz) |  | Baseline (mm) | Current (mm) | Diff (abs) | Diff (%) | Result | Image | Freq (MHz) |  | Baseline (mm) | Current (mm) | Diff (abs) | Diff (%) | Result | Image
2015-08-13 | 6 | Proximal (B1) | 2.61 | 2.57 | 0.04 | 1.35 | PASS | View Image | 6 | Filament 1 | 2.74 | 2.25 | 0.49 | 19.45 | PASS | View Image
|  | Proximal (F1) | 3.23 | 2.1 | 1.13 | 42.56 | FAIL |  |  | Filament 6 | 2.5 | 1.79 | 0.71 | 33.18 | PASS
|  | Distal (B5) | 5.42 | 5.57 | 0.15 | 2.82 | PASS
|  | Distal (F5) | 5.47 | 4.18 | 1.29 | 26.74 | FAIL
Note on I3: Poon, Justin:
Note on Q3: Poon, Justin:

## Area
Area | Known (cm^2) | Measured (cm^2) | Diff (abs) | Diff (%) | Result | Image
2015-07-23 | 3.05 | 2.95 | 0.1 | 3.31 | PASS | View Image
Note on G2: Poon, Justin:

## Volume
Volume | Known (cm^3) | Measured (cm^3) | Diff (abs) | Diff (%) | Result | Images
2015-07-23 | 4 | 4.14 | 0.14 | 3.56 | PASS | View Image 1 | View Image 2 | View Image 3 | View Image 4
2015-07-23 | 20 | 21.45 | 1.45 | 6.99 | FAIL | View Image 1 | View Image 2 | View Image 3 | View Image 4 | View Image 5 | View Image 6 | View Image 7 | View Image 8 | View Image 9
2015-07-23 | 4 | 3.44 | 0.56 | 15.19 | FAIL | View Image 1 | View Image 2
2015-07-23 | 20 | 20.45 | 0.45 | 2.22 | PASS | View Image 1 | View Image 2
Note on G2: Poon, Justin:
Note on H2: Poon, Justin:
Note on I2: Poon, Justin:
Note on J2: Poon, Justin:
Note on G3: Poon, Justin:
Note on H3: Poon, Justin:
Note on I3: Poon, Justin:
Note on J3: Poon, Justin:
Note on K3: Poon, Justin:
Note on L3: Poon, Justin:
Note on M3: Poon, Justin:
Note on N3: Poon, Justin:
Note on O3: Poon, Justin:
Note on G4: Poon, Justin:
Note on H4: Poon, Justin:
Note on G5: Poon, Justin:
Note on H5: Poon, Justin:

## Grid Alignment
Grid Alignment |  | Error (mm) | Result | Image |  | Distance (mm)
2015-07-23 | B1 | 0.7 | PASS | View Image 1 | Probe to G2 | 18.4
| A3 | 2.52 | PASS | View Image 2
| A10 | 1.46 | PASS | View Image 3
| B12 | 2.97 | PASS | View Image 4
| L12 | 1.29 | PASS | View Image 5
| M10 | 0.59 | PASS | View Image 6
| M3 | 1.72 | PASS | View Image 7
| L1 | 1.74 | PASS | View Image 8
| H7 | 0.98 | PASS | View Image 9
| G7 | 1.11 | PASS | View Image 10
| F7 | 2.05 | PASS | View Image 11
| F6 | 0.54 | PASS | View Image 12
| G6 | 0.47 | PASS | View Image 13
| H6 | 1.11 | PASS | View Image 14
Note on E2: Poon, Justin:
Note on E3: Poon, Justin:
Note on E4: Poon, Justin:
Note on E5: Poon, Justin:
Note on E6: Poon, Justin:
Note on E7: Poon, Justin:
Note on E8: Poon, Justin:
Note on E9: Poon, Justin:
Note on E10: Poon, Justin:
Note on E11: Poon, Justin:
Note on E12: Poon, Justin:
Note on E13: Poon, Justin:
Note on E14: Poon, Justin:
Note on E15: Poon, Justin:

## Axial Distance
Axial Distance | Freq (MHz) |  | Known (mm) | Measured (mm) | Diff (abs) | Diff (%) | Result | Image
2015-08-13 | 6 | B1 - B5 | 40 | 40.33 | 0.33 | 0.83 | PASS | View Image
|  | F1 - F5 | 40 | 40.24 | 0.24 | 0.61 | PASS
2015-08-14 | 9 | B1 - B5 | 40 | 40.29 | 0.29 | 0.72 | PASS | View Image
|  | F1 - F5 | 40 | 40.57 | 0.57 | 1.42 | PASS
Note on I2: Poon, Justin:
Note on I5: Poon, Justin:

## Lateral Distance
Lateral Resolution | Axial Plane | [7 empty] | Longitudinal Plane
| Freq (MHz) |  | Known (mm) | Measured (mm) | Diff (abs) | Diff (%) | Result | Image | Freq (MHz) |  | Known (mm) | Measured (mm) | Diff (abs) | Diff (%) | Result | Image
2015-08-14 | 9 | Proximal | 40 | 40.71 | 0.71 | 1.75 | PASS | View Image | 12 | Distance | 40 | 40.01 | 0.01 | 0.03 | PASS | View Image
|  | Distal | 40 | 40.74 | 0.74 | 1.84 | PASS
2015-08-14 | 6 | Proximal | 40 | 40.71 | 0.71 | 1.76 | PASS | View Image | 6 | Distance | 40 | 40.1 | 0.1 | 0.25 | PASS | View Image
|  | Distal | 40 | 40.9 | 0.9 | 2.22 | PASS
Note on I3: Poon, Justin:
Note on Q3: Poon, Justin:
Note on I6: Poon, Justin:
Note on Q6: Poon, Justin: