## Phantom
Phantom | Weight
2015-07-29 | 1500
2015-07-30 | 1600

## Grayscale
Gradient Length | Baseline (mm) | Current (mm) | Result
2015-07-30 | 10.7 | 11.06 | PASS

## Depth
Gradient Length | Axial Plane |  |  | Longitudinal Plane
| Baseline (mm) | Current (mm) | Result | Baseline (mm) | Current (mm) | Result
2015-07-30 | 75.3 | 85.42 | FAIL | 78.8 | 79.2 | PASS
2015-07-30 | 75.3 | 85.42 | FAIL | 78.8 | 79.2 | PASS

## Axial Resolution
Axial Resolution | Axial Plane | [5 empty] | Longitudinal Plane
|  | Baseline (mm) | Current (mm) | Diff (abs) | Diff (%) | Result |  | Baseline (mm) | Current (mm) | Diff (abs) | Diff (%) | Result
2015-07-29 | Proximal (B1) | 0.41 | 0.77 | 0.36 | 60.99 | PASS | Filament 1 | 0.44 | 0.61 | 0.17 | 32.72 | PASS
| Proximal (F1) | 0.55 | 0.51 | 0.04 | 6.79 | PASS | Filament 6 | 0.44 | 0.47 | 0.03 | 6.65 | PASS
| Distal (B5) | 0.69 | 0.73 | 0.04 | 5.86 | PASS
| Distal (F5) | 0.5 | 0.69 | 0.19 | 31.48 | PASS
2015-07-30 | Proximal (B1) | 0.41 | 0.46 | 0.05 | 11.97 | PASS | Filament 1 | 0.44 | 0.5 | 0.06 | 13.36 | PASS
| Proximal (F1) | 0.55 | 0.53 | 0.02 | 3.65 | PASS | Filament 6 | 0.44 | 0.48 | 0.04 | 9.67 | PASS
| Distal (B5) | 0.69 | 0.58 | 0.11 | 17.24 | PASS
| Distal (F5) | 0.5 | 0.62 | 0.12 | 21.39 | PASS

## Lateral Resolution
Lateral Resolution | Axial Plane | [5 empty] | Longitudinal Plane
|  | Baseline (mm) | Current (mm) | Diff (abs) | Diff (%) | Result |  | Baseline (mm) | Current (mm) | Diff (abs) | Diff (%) | Result
2015-07-29 | Proximal (B1) | 1.27 | 2.46 | 1.19 | 63.64 | FAIL | Filament 1 | 1.11 | 2 | 0.89 | 57.09 | PASS
| Proximal (F1) | 1.37 | 2.7 | 1.33 | 65.4 | FAIL | Filament 6 | 1 | 1.76 | 0.76 | 54.96 | PASS
| Distal (B5) | 3.43 | 5.84 | 2.41 | 52.04 | FAIL
| Distal (F5) | 4.3 | 6.24 | 1.94 | 36.78 | FAIL
2015-07-30 | Proximal (B1) | 1.27 | 2.61 | 1.34 | 69.03 | FAIL | Filament 1 | 1.11 | 2.74 | 1.63 | 84.54 | FAIL
| Proximal (F1) | 1.37 | 3.23 | 1.86 | 80.97 | FAIL | Filament 6 | 1 | 2.5 | 1.5 | 85.85 | FAIL
| Distal (B5) | 3.43 | 5.42 | 1.99 | 45.03 | FAIL
| Distal (F5) | 4.3 | 5.47 | 1.17 | 23.91 | FAIL

## Axial Distance
Axial Distance |  | Known (mm) | Measured (mm) | Diff (abs) | Diff (%) | Result
2015-07-29 | B1 - B5 | 40 | 40.57 | 0.57 | 1.41 | PASS
| F1 - F5 | 40 | 39.93 | 0.07 | 0.17 | PASS
2015-07-30 | B1 - B5 | 40 | 39.96 | 0.04 | 0.09 | PASS
| F1 - F5 | 40 | 39.81 | 0.19 | 0.47 | PASS